## Slide 1
ALAGA: A Low-Cost Soft Inflatable Healthcare Companion Robot with Non-Contact Temperature, Pulse Oximetry, and Pain-Scale Triage for Evacuation Centers
[Surname, Firstname M.]
[Surname, Firstname M.]
[Surname, Firstname M.]
Notes: [Researcher 1] | 0:00 to 0:15
Good morning to the panel. We are [names], with our adviser, [adviser]. Our study is titled ALAGA, a low-cost soft inflatable healthcare companion robot.

## Slide 2
INTRODUCTION
Background of the Study
2
• Mayon was raised to Alert Level 3 in January 2026, and ash fell on 87 villages in May (DZRH, 2026).
• Evacuation centers fill up, and ash irritates the eyes and airways (Horwell & Baxter, 2006).
• A few health workers must check every evacuee's temperature, pulse, oxygen, and pain.
• Healthcare robots exist, but they are rigid and cost more than a barangay can pay.
Table 1. The problem in figures
Indicator | Value | Source
Mayon alert level, Jan 2026 | Level 3 | DZRH (2026)
Villages under ash, May 2026 | 87 | Khaosod English (2026)
Families evacuated, May 2026 | 300+ | Khaosod English (2026)
SpO₂ accuracy limit (Arms) | ≤ 4% | ISO (2017b)
Pain scale used in screening | 0–10 NRS | Hawker et al. (2011)
Purpose of the study. To find out whether a low-cost soft robot, inspired by the healthcare companion of Big Hero 6, can screen evacuees accurately.
Notes: [Researcher 1] | 0:15 to 0:40
Mayon was raised to Alert Level 3 in January 2026, and in May ash fell on 87 villages. More than 300 families moved to evacuation centers. There, a few health workers must check every evacuee for fever, low oxygen, and pain. Healthcare robots exist, but they are rigid and costly. Our purpose: to find out whether a low-cost soft robot, inspired by the healthcare companion in Big Hero 6, can screen evacuees accurately.

## Slide 3
RRL/RRS
Related Literature
3
Table 2. Related literature and how the study uses it
Area | Key sources | Use in the present study
Soft and inflatable robots | Rus & Tolley (2015); Sanan et al. (2011) | Air-filled body that is safe to touch
Acceptance of care robots | Broadbent et al. (2009); Mori et al. (2012) | Friendly look; simple two-dot face
Vital-sign sensing | Jubran (2015); Allen (2007); ISO (2017a, 2017b) | SpO₂, pulse, and IR temperature limits
Method comparison, usability | Bland & Altman (1986); Brooke (1996) | Agreement statistics and SUS
Synthesis. Soft bodies are safer and friendlier, and low-cost sensors can screen vital signs if checked against a reference.
Notes: [Researcher 1] | 0:40 to 0:58
From the literature: soft and inflatable robots are safer to touch; older adults accept robots that look friendly and are easy to use; pulse oximetry and infrared thermometry can screen vital signs; and Bland and Altman's limits of agreement test whether a new instrument can be trusted.

## Slide 4
RRL/RRS
Related Studies
4
Table 3. Related studies and their relation to the present study
Study | What was done | Relation to the present study
Sanan et al. (2011) | Built an inflatable arm for safe contact with people. | Supports an air-filled body.
Broadbent et al. (2009) | Reviewed what makes older adults accept care robots. | Guides the friendly design.
Bland & Altman (1986) | Proposed limits of agreement for two methods. | Method used to test ALAGA.
Synthesis. No low-cost robot joins a soft body, vital-sign screening, and pain triage for evacuation centers.
Notes: [Researcher 1] | 0:58 to 1:15
Related studies built an inflatable arm for safe contact, reviewed acceptance of care robots, and gave us the agreement method. None joins a soft body, vital-sign screening, and pain triage for evacuation centers.

## Slide 5
THEORETICAL/CONCEPTUAL FRAMEWORK PARADIGM
Theoretical Framework
5
Research gap
Theory or concept that addresses it
G1 Few health workers
Hundreds of evacuees need routine checks.
Socially assistive robot
Robot screens; health workers treat.
G2 Robots can scare
Rigid or almost-human robots feel unsafe.
Soft body + plain face
Air body; simple face (Mori et al., 2012).
G3 Cheap sensors vary
Low-cost modules are rarely validated.
PPG, IR + agreement
Calibrate, then Bland–Altman test.
G4 Who needs help first
Raw numbers mean little to evacuees.
State machine + limits
GREEN, YELLOW, or RED with spoken advice.
Figure 1. Research gaps (left) and the theories and concepts used to address them (right)
Notes: [Researcher 2] | 1:15 to 1:35
Our framework answers four gaps. Few health workers: a socially assistive robot does the routine checks. Robots can scare people: an air-filled body and a plain two-dot face. Cheap sensors vary: we calibrate and test agreement. Who needs help first: a state machine gives GREEN, YELLOW, or RED.

## Slide 6
THEORETICAL/CONCEPTUAL FRAMEWORK PARADIGM
Conceptual Paradigm
6
INPUT
• Ripstop nylon body and 12 V blower fan
• ESP32, MLX90614, MAX30102, HC-SR04
• OLED, DFPlayer, speaker, pain dial, button
• Reference thermometer and pulse oximeter
• Consenting volunteers (IRB-approved)
PROCESS
1. Sew and seal the inflatable body
2. Wire sensors; program the state machine
3. Calibrate on 10 volunteers
4. Screen volunteers vs reference
5. Logic tests; soft-vs-rigid usability
6. Bland–Altman, t-test, Wilcoxon, cost
OUTPUT
• ALAGA prototype
• Bias, limits of agreement, Arms
• Flag sensitivity, specificity
• Screening time (s)
• SUS and approachability
• Cost per unit and per day
Feedback: results tune the calibration offsets and flag limits
Figure 2. Conceptual paradigm of the study
Notes: [Researcher 2] | 1:35 to 1:58
Our paradigm follows input, process, output. Inputs are the body, sensors, reference instruments, and volunteers. The process is building, calibration, trials, and analysis. Outputs answer each objective, and results feed back to tune the offsets and limits.

## Slide 7
SIGNIFICANCE OF THE STUDY
7
Evacuees. Children and older adults are screened by a friendly, soft robot without a long wait.
Health workers. Routine checks are shared, so more time goes to evacuees flagged YELLOW or RED.
LGUs and the PDRRMO. An affordable screening aid for every evacuation center.
Schools used as shelters. The same robot serves the school clinic between disasters.
Future researchers. A tested open design and a way to validate low-cost sensors.
The study supports SDG 3 (good health), SDG 9 (innovation), and SDG 11.5 (fewer people affected by disasters).
Notes: [Researcher 2] | 1:58 to 2:15
Evacuees get a friendly, quick screening. Health workers spend more time on flagged evacuees. LGUs get an affordable aid for every center, schools a clinic helper, and future researchers a tested open design. The study supports SDGs 3, 9, and 11.5.

## Slide 8
OBJECTIVES OF THE STUDY / SOP
8
General objective. This study aimed to design, build, and evaluate ALAGA, a low-cost soft inflatable healthcare companion robot that screens temperature, pulse rate, SpO₂, and pain and flags evacuees who need a health worker.
Specifically, it seeks to answer the following questions:
1. How closely do ALAGA's temperature, pulse rate, and SpO₂ readings agree with reference instruments (bias, limits of agreement, MAE, Arms)?
2. How accurately does ALAGA flag readings outside the screening limits (sensitivity, specificity, accuracy)?
3. How long does ALAGA take to complete one screening, from wake-up to spoken result?
4. How do users rate ALAGA's usability, and is there a significant difference in approachability between the soft body and a rigid shell?
5. How much does ALAGA cost to build and to run per day?
Notes: [Researcher 2] | 2:15 to 2:35
Our general objective: to design, build, and evaluate ALAGA. Specifically: agreement with reference instruments, accuracy of flags, screening time, usability and approachability of the soft body versus a rigid shell, and cost.

## Slide 9
HYPOTHESES
Null Hypotheses
9
The hypotheses were stated before testing and are tested at the 0.05 level of significance.
Code | Hypothesis | Test or decision rule
H₀1 | There is no significant difference between ALAGA's readings and the reference readings for temperature, pulse rate, and SpO₂. | Paired t-test; Bland–Altman bias and 95% limits of agreement
H₀2 | There is no significant difference in approachability ratings between the soft inflatable body and the rigid shell. | Wilcoxon signed-rank test
Reject H₀ when p < 0.05.
Notes: [Researcher 3] | 2:35 to 2:50
H₀1: there is no significant difference between ALAGA's readings and the reference readings; we use a paired t-test and Bland–Altman analysis. H₀2: there is no significant difference in approachability between the soft body and a rigid shell; we use the Wilcoxon signed-rank test. We reject when p is below 0.05.

## Slide 10
SCOPE AND DELIMITATION
10
Scope
• One stationary robot, about 1.1 m tall, fan-inflated
• Forehead temperature, pulse rate, SpO₂, pain 0–10
• Filipino and English prompts; chest display; log
• Healthy volunteers aged 15+, after IRB approval
• Water-bath bench test of the IR thermometer
Delimitation
• No walking, arms, or touching
• Screens and refers only; not a medical device
• No blood pressure or respiratory rate
• Fever and low SpO₂ simulated, not from patients
• Tested in school, not in a real evacuation center
Table 4. Trials Conducted in the Study
Trial set | Measures | Number | Reference
Bench | IR temperature, 35–40 °C target | 6 × 10 | Probe thermometer
Calibration | Temperature, PR, SpO₂ | 10 volunteers | Clinic instruments
Validation | Temperature, PR, SpO₂, time | ____ volunteers | Clinic instruments
Logic + usability | Flags; SUS; approachability | 60 cases; ____ users | Limits; soft vs rigid
Notes: [Researcher 3] | 2:50 to 3:05
We cover one stationary robot that screens temperature, pulse, oxygen, and pain, tested on healthy volunteers after IRB approval. ALAGA does not walk or diagnose, and fever and low oxygen are simulated by a water bath and logic test cases. Table 4 lists our trials.

## Slide 11
METHODOLOGY
Research Design
11
This is a quantitative developmental and method-comparison study. The robot is built and calibrated first; then its readings are compared with reference instruments, and users rate it.
1. Build and calibrate
2. Program logic
3. Screen volunteers
4. Analyze results
Table 5. Variables of the study
Independent (varied) | Dependent (measured) | Controlled (held fixed)
Measurement method (ALAGA; reference) Target temperature (°C) Body type (soft; rigid) | Bias, LoA, MAE, Arms Sensitivity, specificity Screening time; SUS; approachability | Room 26–30 °C; 5-min rest Forehead 3–5 cm; left index finger Same prompts; order alternated
Notes: [Researcher 3] | 3:05 to 3:22
This is a quantitative developmental and method-comparison study. We build and calibrate, program the logic, screen volunteers, and analyze. We vary the measurement method, target temperature, and body type, and we measure agreement, flag accuracy, time, and ratings.

## Slide 12
METHODOLOGY
Materials and Device
12
Device specification
• Body: white ripstop nylon, about 1.1 m tall, kept inflated by a 12 V blower fan in the base.
• Chest panel: MLX90614 and HC-SR04 at forehead height; MAX30102 finger cradle.
• Controls: 0–10 pain dial, chest button, and a sound sensor that wakes ALAGA.
• Voice and display: DFPlayer Mini with 3 W speaker; 0.96-in OLED on the chest.
• Controller and power: ESP32 state machine; 12 V 7 Ah battery, 5 V regulator; CSV log.
Table 6. Materials and Components of ALAGA
Material / Component | Quantity | Purpose
White 70D ripstop nylon, seam tape, zipper | 3 m; 1 set | Inflatable body and head
12 V DC blower fan with finger guard | 1 | Keeps the body inflated
ESP32 development board | 1 | Controller
MLX90614 (GY-906) infrared thermometer | 1 | Forehead temperature
MAX30102 pulse oximeter module | 1 | Pulse rate and SpO₂
HC-SR04 ultrasonic sensor | 1 | Forehead distance check
0.96-in OLED display (SSD1306) | 1 | Chest display
DFPlayer Mini, microSD card, 3 W speaker | 1 set | Spoken prompts
KY-037 sound sensor, push button, 10 kΩ dial | 1 each | Wake-up and pain score
12 V 7 Ah SLA battery, LM2596 regulator, fuse | 1 set | Power
IRF520 MOSFET module, wires, perfboard | 1 set | Fan switch and wiring
Reference forehead thermometer and pulse oximeter | 1 each | Reference readings (borrowed)
Notes: [Researcher 3] | 3:22 to 3:40
ALAGA's body is white ripstop nylon kept full by a 12-volt fan. The chest panel holds the infrared thermometer, distance sensor, finger cradle, pain dial, and display. An ESP32 runs everything from a 12-volt battery. The parts are estimated at about 4,600 pesos.

## Slide 13
METHODOLOGY
Data Gathering Procedure
13
Step 1
Secure IRB approval, consent
Step 2
Sew and seal the body
Step 3
Wire sensors, audio, fan
Step 4
Program state machine
Step 5
Bench-test IR in water bath
Step 6
Calibrate on 10 volunteers
Step 7
Screen validation volunteers
Step 8
Run 60 logic test cases
Step 9
Usability: soft vs rigid
Step 10
Bland–Altman, t-test, Wilcoxon
Figure 3. Data gathering procedure
Notes: [Researcher 1] | 3:40 to 4:00
The procedure has ten steps: IRB approval and consent; sew the body; wire the sensors; program the state machine; bench-test with a water bath; calibrate on ten volunteers; screen validation volunteers; run logic tests; run the usability test; and analyze.

## Slide 14
METHODOLOGY
Treatment of Data
14
Formulas
d = ALAGA − reference
difference per paired reading
Bias = Σd ÷ n; LoA = Bias ± 1.96 SD
mean bias, 95% limits of agreement
Arms = √(Σd² ÷ n)
SpO₂ root-mean-square error (≤ 4%)
Sensitivity = TP ÷ (TP + FN)
flags correctly raised
Specificity = TN ÷ (TN + FP)
normal results correctly passed
Statistical treatment
SOP | Question | Tool
S1 | Agreement with reference instruments | Bias, LoA, MAE, Arms; paired t-test
S2 | Accuracy of flags | Confusion matrix; sensitivity, specificity
S3 | Screening time | Mean ± SD
S4 | Usability and approachability | SUS; Wilcoxon signed-rank
S5 | Cost to build and run | Itemized cost; PHP per day
α = 0.05; software: Microsoft Excel and JASP.
Notes: [Researcher 1] | 4:00 to 4:20
For each paired reading we take the difference. Bias is the mean difference, and the limits of agreement are bias plus or minus 1.96 standard deviations. Arms is the SpO₂ root-mean-square error, accepted at 4 percent or less. Flags are judged by sensitivity and specificity. We use Excel and JASP at alpha 0.05.

## Slide 15
METHODOLOGY
Ethical Considerations and Safety
15
Ethics.
Human participants: plan approved by the school IRB before testing (ISEF Form 4).
Informed consent; assent and parent consent for minors; may stop anytime.
Non-invasive only; codes, not names; flagged readings referred to the nurse.
ALAGA screens and refers; it gives no diagnosis. No data were fabricated.
Risk and safety.
Blower fan: metal finger guard; fan sealed inside the base.
12 V battery: inline fuse; wiring in a closed box; supervised charging.
Shared sensors: wiped with 70% alcohol between volunteers.
Wrong reading: reference taken for every volunteer; nurse referral.
Personal data: password-protected and deleted after the fair.
Notes: [Researcher 1] | 4:20 to 4:35
Because we work with people, the school IRB approves our plan first, and every volunteer gives consent. Tests are non-invasive, data are coded, and any abnormal reading goes to the nurse. The fan is guarded, the battery is fused, and sensors are wiped with alcohol.

## Slide 16
QUESTIONS
& ANSWERS
Notes: [Researcher 3] | 4:35 to 4:42
That ends our presentation. We are now ready for your questions. Thank you.

## Slide 17
THANK YOU
Notes: Leave this slide up after the closing line of slide 17.