ENGINEERING NOTEBOOK

KALINGA

Development of a 3D-Printed, Baymax-Inspired Wearable Health Pendant

The pendant to be built (front view).

Team: ______, ______, ______     Adviser: ______     School: ______

How to keep this notebook. Write every entry on the day the work is done. Sign and date each entry and have a witness sign. Do not erase or tear pages: cross out mistakes with one line. Paste printouts, photos, and screenshots with the date. Record what failed as carefully as what worked; judges look for the design process, not only the final product.

Entry 1: Problem and Design Criteria

Date: ______     Recorded by: ______     Witnessed by: ______

Problem. Older adults and people with chronic illness are often alone at home while their families work. Falls go unnoticed, doses are missed, and fevers rise unseen. Existing products each do one job and cost more together.

Goal. Build one low-cost wearable that checks vital signs, presents each dose on time, alerts the family, and shows everything in a phone app.

| Criterion | Target | How it will be checked |
| --- | --- | --- |
| Cost | Below PHP 2,500 for parts | Receipts (Appendix A.3) |
| Size and weight | Fits in a hand; about 60 × 83 × 39 mm; under 120 g | Ruler and kitchen scale |
| Temperature accuracy | Bias within ±0.3 °C after calibration | Bland–Altman vs clinic thermometer |
| SpO₂ accuracy | Arms ≤ 4% (ISO 80601-2-61) | Bland–Altman vs reference oximeter |
| Fall detection | Sensitivity and specificity ≥ 90% | Mat falls and daily activities |
| Alert time | Under 60 s to the app | Stopwatch and app time stamps |
| Doses | Correct slot every time; within 1 min of schedule | 5-day dose log |
| Battery | At least 16 h per charge | Battery log |
| Ease of use | SUS above 68 | Questionnaire |

Entry 2: Background Research

Date: ______     Recorded by: ______     Witnessed by: ______

•	Pulse oximetry: red and infrared light through the fingertip give SpO₂ and pulse rate (Allen, 2007; Jubran, 2015). Readings drop in accuracy with movement, cold fingers, nail polish, and darker skin.

•	Infrared thermometry: the MLX90614 reads skin temperature without contact; forehead skin is cooler than the body core, so a calibration equation is needed (Fraden, 2016).

•	Falls: free fall (< 0.5 g), then impact (> 2.5 g), then stillness is a simple, proven rule (Bourke et al., 2007).

•	Medicine: about half of long-term patients do not take medicines as prescribed; timed reminders help (WHO, 2003).

•	Data: the Firebase Realtime Database has a REST interface that a microcontroller can call over HTTPS and a web app can stream live (Google, n.d.). Health data must be protected (Republic Act No. 10173, 2012).

Sources read by the team and notes: ______

Entry 3: Ideas and Decision Matrix

Date: ______     Recorded by: ______     Witnessed by: ______

Each idea is scored 1 (poor) to 5 (best) on each criterion and multiplied by the weight. The scores below are the starting estimates; change them after your own tests and sign the decision.

| Option for each function | Cost ×3 | Size ×2 | Ease to build ×2 | Reliability ×3 | Total |
| --- | --- | --- | --- | --- | --- |
| Dispensing: 4-slot carousel + SG90 servo | 5 | 3 | 4 | 4 | 41 |
| Dispensing: single lid with servo latch | 5 | 4 | 4 | 3 | 40 |
| Dispensing: sliding drawer + linear servo | 3 | 4 | 2 | 3 | 30 |
| Temperature: MLX90614 non-contact | 3 | 4 | 5 | 4 | 39 |
| Temperature: MAX30205 skin contact | 4 | 5 | 3 | 3 | 37 |
| Link: Wi-Fi + Firebase + web app | 5 | 5 | 3 | 4 | 43 |
| Link: Bluetooth to the older adult's phone | 5 | 5 | 3 | 2 | 37 |
| Link: GSM SMS module | 3 | 2 | 3 | 3 | 28 |

Chosen design: carousel + SG90, MLX90614, Wi-Fi + Firebase.   Team signatures: ______

Entry 4: Bill of Materials and Tools

Date: ______     Recorded by: ______     Witnessed by: ______

| Item | Qty | Est. PHP |
| --- | --- | --- |
| ESP32-C3 SuperMini | 1 | 220 |
| MLX90614 (GY-906) module | 1 | 550 |
| MAX30102 module | 1 | 180 |
| MPU6050 (GY-521) module | 1 | 120 |
| 0.91-in OLED, 128 × 32, I²C | 1 | 150 |
| SG90 micro servo | 1 | 120 |
| LiPo 3.7 V 1000 mAh (503450) | 1 | 220 |
| TP4056 USB-C charger with protection | 1 | 45 |
| Coin vibration motor, 3 V active buzzer | 1 set | 50 |
| AO3400, S8050, 1N4148, resistors | 1 set | 30 |
| Tactile buttons with caps, slide switch | 1 set | 35 |
| Perfboard, wires, headers | 1 lot | 80 |
| PLA filament or print service (about 60 g) | 1 lot | 250 |
| Black acrylic paint, sandpaper | 1 lot | 40 |
| Breakaway lanyard, M2 screws | 1 set | 60 |
| TOTAL (estimated) |  | 2,150 |

Tools. 3D printer (FDM, 0.4-mm nozzle) or a print shop; slicer (UltiMaker Cura, PrusaSlicer, or Bambu Studio); soldering iron, solder, flux; wire stripper; multimeter; hot-glue gun; small files and sandpaper (220 and 400); flush cutters; laptop with Arduino IDE 2; USB-C cable; placebo candy tablets for testing.

Entry 5: CAD Design of the Shell

Date: ______     Recorded by: ______     Witnessed by: ______

The shell is an egg-shaped body (two spheres joined and flattened) split into a front and a back shell, plus a 36-mm pill carousel. The model is a Python script, so every size can be changed and the STL files regenerated:

pip install manifold3d trimesh      then      python cad/kalinga_pendant.py

| Feature | Size (mm) | Why |
| --- | --- | --- |
| Wall thickness | 2.0 | Strong enough for PLA; light |
| Overall | 60 × 83 × 39 | Fits the SG90 servo and carousel |
| OLED window | 24 × 8 | 0.91-in OLED active area 22.4 × 5.6 |
| Finger pad window | 13 × 13 | MAX30102 module |
| Thermometer window (top) | Ø 9.8 | MLX90614 TO-39 can, Ø 9.2 |
| Pill window | Ø 11 | Pocket Ø 9.6 under it |
| Carousel | Ø 36 × 7, 4 pockets | Pockets at 45°, 90°, 135°, 180°; solid at 0° = closed |
| Button holes | Ø 6.6 | 6 × 6 mm tactile button caps |
| USB-C and switch slots | 10 × 4, 9 × 4 | TP4056 port, slide switch |
| Lip | 3 deep, 1.2 thick | Locates the front shell on the back shell |

Front shell rendered from front_shell.stl.

Exploded view: back shell with servo cradle, pill carousel, and front shell.

Changes made to the CAD and why: ______

Entry 6: 3D Printing

Date: ______     Recorded by: ______     Witnessed by: ______

| Setting | Shells | Carousel |
| --- | --- | --- |
| Material | White PLA | White or red PLA |
| Nozzle / layer height | 0.4 mm / 0.2 mm | 0.4 mm / 0.2 mm |
| Walls / top and bottom layers | 3 / 4 | 3 / 4 |
| Infill | 20% gyroid | 30% gyroid |
| Nozzle / bed temperature | 205 °C / 60 °C | 205 °C / 60 °C |
| Orientation | Flat split edge on the bed (open side down) | Flat, pockets facing up |
| Supports | Tree supports, everywhere; 5-mm brim | None |
| Estimated time and filament (from slicer) | ______ h, ______ g | ______ h, ______ g |

1.	Print the carousel first as a test: check that a placebo tablet fits each pocket and that the SG90 horn sits flat in its recess.

2.	Print the back shell, then the front shell. Watch the first layer; stop and level the bed if lines do not stick.

3.	Remove supports with flush cutters and pliers. Sand edges with 220 then 400 grit.

4.	Dry-fit the two shells: the lip must slide in with light pressure. If tight, sand the lip; if loose, add one layer of tape.

5.	Paint the two eye recesses with black acrylic; let dry for 1 hour. Keep the face design original (two dots, no copied logos) so it can be filed as an Industrial Design.

6.	Weigh and measure each part and paste photos here.

| Part | Print date | Time (h) | Mass (g) | Problems seen | Fix |
| --- | --- | --- | --- | --- | --- |

Entry 7: Electronics Assembly

Date: ______     Recorded by: ______     Witnessed by: ______

Wiring diagram.

| Part | Connects to (ESP32-C3) |
| --- | --- |
| MLX90614, MAX30102, MPU6050, OLED | GPIO 6 / 7 — I²C at 100 kHz, 3.3 V |
| SG90 servo | GPIO 5 / GPIO 10 — AO3400 switches the servo ground |
| OK and SOS buttons | GPIO 2 / GPIO 1 — Internal pull-ups |
| Buzzer; vibration motor | GPIO 3; GPIO 21 — 1 kΩ base resistor; 1N4148 |
| Battery divider | GPIO 4 — Reads 0–4.2 V |
| TP4056 OUT+ / OUT– | 5V / GND — LiPo on B+ / B– |

1.	Solder headers on each module. Test each module alone with the ESP32-C3 and an I²C scan: expect 0x5A (MLX90614), 0x57 (MAX30102), 0x68 (MPU6050), and 0x3C (OLED).

2.	Build the small parts on perfboard: S8050 with 1 kΩ base resistor and 1N4148 across the vibration motor; AO3400 with 10 kΩ gate-to-ground resistor for the servo; 100 kΩ/100 kΩ battery divider to GPIO 4.

3.	LiPo safety: connect the battery to the TP4056 B+ and B– only, with the correct polarity, and never short, bend, or puncture it. Charge only under supervision.

4.	Wire TP4056 OUT+ through the slide switch to the ESP32-C3 5V pin and the servo red wire; OUT– to GND.

5.	Measure with the multimeter before powering: battery 3.7–4.2 V; no short between 5V and GND.

6.	Record each test below.

| Module | Test | Expected | Result | OK? |
| --- | --- | --- | --- | --- |

Entry 8: Programming the Pendant

Date: ______     Recorded by: ______     Witnessed by: ______

1.	Install Arduino IDE 2 and “esp32 by Espressif Systems” in Boards Manager.

2.	Install libraries: Adafruit MLX90614, Adafruit SSD1306, Adafruit GFX, SparkFun MAX3010x, ESP32Servo, ArduinoJson.

3.	Open firmware/kalinga_pendant/kalinga_pendant.ino. Type the Wi-Fi name and password, the Firebase Web API key and database URL, the pendant's account email and password, and the patient id.

4.	Board: ESP32C3 Dev Module; USB CDC On Boot: Enabled. Upload. Open Serial Monitor at 115200.

5.	Check: the OLED shows the time and battery; the carousel moves to the closed position; holding SOS for 2 s sends an SOS alert.

Entry 9: Firebase and the KALINGA Monitor App

Date: ______     Recorded by: ______     Witnessed by: ______

How the pendant, the database, and the app connect.

1.	Go to console.firebase.google.com, create a project, and add a Web app. Copy the Web API key.

2.	Authentication → Sign-in method → enable Email/Password. Add one user for the pendant and one for each family member. Copy each user's UID.

3.	Realtime Database → Create (location: Singapore, asia-southeast1) → start in locked mode. Copy the database URL.

4.	Rules: paste app/database.rules.json and publish.

5.	Data: add members → <patient id> → <each UID> = true (for the pendant and every family member).

6.	Host the app folder on GitHub Pages or Firebase Hosting (HTTPS is needed to install it). Open it on a phone, add it to the home screen, enter the key, URL, patient id, email, and password, and sign in.

7.	Use “Try the demo” to show the app to judges without the pendant.

KALINGA Monitor screens (demo mode).

Entry 10: Final Assembly

Date: ______     Recorded by: ______     Witnessed by: ______

1.	Set the servo to the closed position (power on once; the program moves it to 0°) before fitting the horn and carousel, so the solid sector sits under the window.

2.	Fix the servo in the back-shell cradle with a drop of hot glue; screw the single-arm horn into the carousel recess.

3.	Hot-glue the OLED behind its window, the MAX30102 behind the finger window (sensor flush with the surface), and the MLX90614 can in the top hole.

4.	Place the ESP32-C3, MPU6050 (flat, near the center), battery, and TP4056 with double-sided foam tape; keep the USB-C port at its slot.

5.	Fit the buttons and caps, close the shells, and add two M2 screws or a drop of glue at the bottom.

6.	Thread the breakaway lanyard. Fill the four pockets with placebo tablets and run one dose cycle.

Entry 11: Calibration

Date: ______     Recorded by: ______     Witnessed by: ______

1.	Bench test of the thermometer with a black-painted can in a water bath at 35.0, 36.0, 37.0, 37.5, 38.0, and 40.0 °C against a probe thermometer.

2.	Calibration set (10 volunteers, after IRB approval): TEMP_A = SLOPE(reference, raw), TEMP_B = INTERCEPT(reference, raw); PR and SpO₂ offsets = mean of (reference − pendant).

3.	Fall thresholds: 10 mat falls and 10 daily activities by the researchers; adjust FREE_FALL_G and IMPACT_G, then keep them fixed.

| Target or volunteer | Reference | Pendant raw | Pendant calibrated | Difference |
| --- | --- | --- | --- | --- |

Entry 12: Test Records

Date: ______     Recorded by: ______     Witnessed by: ______

Falls and daily activities

| Trial | Event | Detected? | Cancelled? | Peak g | Alert time (s) |
| --- | --- | --- | --- | --- | --- |

Doses

| Date | Slot | Scheduled | Opened at | Recorded as | Observer says |
| --- | --- | --- | --- | --- | --- |

Battery

| Charge no. | Full at | Empty at | Hours | Notes |
| --- | --- | --- | --- | --- |

Entry 13: Problems and Iterations

Date: ______     Recorded by: ______     Witnessed by: ______

Write each problem the day it happens, what caused it, what you changed, and what happened after.

| Date | Problem | Cause | Change made | Result |
| --- | --- | --- | --- | --- |

Entry 14: Safety Checklist and Reflection

Date: ______     Recorded by: ______     Witnessed by: ______

•	Adult supervision for the 3D printer, soldering, and LiPo charging.

•	IRB approval and signed consent before any test with people; placebo tablets only.

•	Falls simulated only by the researchers on 20-cm mats with pads.

•	Breakaway lanyard so the cord cannot strangle the wearer.

•	Firebase locked by rules; coded patient id; data deleted after the fair.

What worked: ______

What we would change next: ______