## Slide 1
KALINGA: A Low-Cost Baymax-Inspired Wearable Health Pendant with Vital-Sign Checks, Fall Alerts, Timed Pill Dispensing, and a Family Monitoring App
[Surname, Firstname M.]
[Surname, Firstname M.]
[Surname, Firstname M.]
Notes: [Researcher 1] | 0:00 to 0:15
Good morning to the panel. We are [names], with our adviser, [adviser]. Our entry is KALINGA, a low-cost Baymax-inspired wearable health pendant with a family monitoring app.

## Slide 2
INTRODUCTION
Background of the Study
2
• Many older adults and people with chronic illness are alone at home while their families work.
• Adults over 60 suffer the most fatal falls, and help often comes late (WHO, 2021).
• About half of patients on long-term medicines miss or mistime doses (WHO, 2003).
• Smartwatches, alert pendants, and pill dispensers each do one job and cost more together.
Table 1. The problem in figures
Indicator | Value | Source
Group with the most fatal falls | Adults 60+ | WHO (2021)
Adherence to long-term medicines | About 50% | WHO (2003)
SpO₂ accuracy limit (Arms) | ≤ 4% | ISO (2017b)
Average SUS score of products | 68 | Sauro (2011)
Estimated cost of KALINGA | PHP 2,150 | Appendix A.3
Purpose of the study. To develop one low-cost Baymax-inspired pendant that checks vital signs, times medicine, and alerts the family through an app.
Notes: [Researcher 1] | 0:15 to 0:40
Many older adults are alone at home while their families work. Falls go unnoticed, doses are missed, and fevers rise unseen. The products that help each do one job and cost more together. Our purpose: one affordable pendant that checks, reminds, and alerts the family through an app.

## Slide 3
RRL/RRS
Related Literature
3
Table 2. Related literature and how the study uses it
Area | Key sources | Use in the present study
Pulse oximetry and IR thermometry | Allen (2007); Jubran (2015); Fraden (2016) | How the pendant measures
Fall detection | Bourke et al. (2007); WHO (2021) | Free fall, impact, stillness rule
Medicine adherence | WHO (2003) | Timed doses and reminders
Agreement and usability | Bland & Altman (1986); Brooke (1996); Sauro (2011) | Tests of accuracy and SUS
Synthesis. Low-cost sensors can check vital signs and detect falls, and timed doses improve adherence.
Notes: [Researcher 1] | 0:40 to 0:58
The literature tells us how pulse oximeters and infrared thermometers work, how falls can be detected with an accelerometer, why timed doses help, and how to test accuracy and usability.

## Slide 4
RRL/RRS
Related Studies
4
Table 3. Related studies and their relation to the present study
Study | What was done | Relation to the present study
Bourke et al. (2007) | Detected falls with accelerometer thresholds. | Basis of the fall rule.
Jubran (2015) | Reviewed pulse oximetry and its errors. | Sets SpO₂ limits and cautions.
Bland & Altman (1986) | Proposed limits of agreement for two methods. | Method used to test KALINGA.
Synthesis. No low-cost wearable joins vital checks, fall alerts, timed doses, and a family app.
Notes: [Researcher 1] | 0:58 to 1:15
Related studies gave us the fall rule, the limits of pulse oximetry, and the agreement method. None joins all of these in one low-cost wearable.

## Slide 5
THEORETICAL/CONCEPTUAL FRAMEWORK PARADIGM
Theoretical Framework
5
Research gap
Theory or concept that addresses it
G1 Help comes late
Falls go unnoticed for hours.
Threshold fall detection
Alert reaches the app in seconds.
G2 Missed doses
Medicines are forgotten or mistimed.
Timed pill carousel
Opens the right slot on time.
G3 Cheap sensors vary
Low-cost modules are rarely validated.
PPG, IR + agreement
Calibrate, then Bland–Altman.
G4 Family far away
No one sees the readings.
IoT + cloud database
Live app for the family.
Figure 1. Research gaps (left) and the theories and concepts used to address them (right)
Notes: [Researcher 2] | 1:15 to 1:35
Four gaps: help comes late, so the pendant alerts in seconds; doses are missed, so a carousel opens the right slot on time; cheap sensors vary, so we calibrate and test; the family is far, so a live app shows everything.

## Slide 6
THEORETICAL/CONCEPTUAL FRAMEWORK PARADIGM
Conceptual Paradigm
6
INPUT
• ESP32-C3, MLX90614, MAX30102, MPU6050
• OLED, SG90 servo, buzzer, vibration
• LiPo battery, USB-C charger
• White PLA, 3D printer
• Firebase, app; reference instruments
PROCESS
1. Design and 3D-print the shell
2. Wire and program the pendant
3. Set up Firebase and the app
4. Calibrate sensors and fall rule
5. Agreement, fall, dose, battery tests
6. Usability; statistics
OUTPUT
• KALINGA pendant and app
• Bias, limits of agreement, Arms
• Fall sensitivity, specificity
• Correct-slot and on-time doses
• Delivery rate, battery life, SUS
• Cost comparison
Feedback: results tune the calibration, fall thresholds, and reminders
Figure 2. Conceptual paradigm of the study
Notes: [Researcher 2] | 1:35 to 1:58
Inputs are the sensors, the servo, the battery, the 3D printer, and Firebase. We print, wire, program, calibrate, and test. Outputs answer each objective, and the results tune the pendant.

## Slide 7
SIGNIFICANCE OF THE STUDY
7
Older adults and patients. Checks, medicine timing, and help at the press of a button.
Working families and OFWs. Live readings and alerts on their phones, anywhere.
Barangay health workers. A low-cost way to watch seniors who live alone.
Innovators and LGUs. A local, 3D-printable design that can be repaired and improved.
Future researchers. An open design and a tested method for low-cost health sensors.
The study supports SDG 3 (good health), SDG 9 (innovation), and SDG 10 (reduced inequalities).
Notes: [Researcher 2] | 1:58 to 2:15
Older adults get checks and help at a button press, families see everything from work, and health workers get a low-cost aid. The study supports SDGs 3, 9, and 10.

## Slide 8
OBJECTIVES OF THE STUDY / SOP
8
General objective. This study aimed to develop and evaluate KALINGA, a low-cost Baymax-inspired wearable health pendant with vital-sign checks, fall alerts, timed pill dispensing, and a family monitoring app.
Specifically, it seeks to answer the following questions:
1. How closely do its temperature, pulse rate, and SpO₂ agree with reference instruments?
2. How accurately does it detect falls, and how fast do alerts reach the app?
3. How accurately and on time does the carousel present each dose?
4. What share of data reaches the app, and how long does the battery last?
5. How do older adults and caregivers rate its usability (SUS)?
6. How does its cost compare with commercial products?
Notes: [Researcher 2] | 2:15 to 2:35
Our general objective: develop and evaluate KALINGA. Specifically: accuracy, falls and alerts, doses, data delivery and battery, usability, and cost.

## Slide 9
HYPOTHESES
Null Hypotheses
9
The hypotheses were stated before testing and are tested at the 0.05 level of significance.
Code | Hypothesis | Test or decision rule
H₀1 | There is no significant difference between KALINGA's readings and the reference readings. | Paired t-test; Bland–Altman
H₀2 | KALINGA's fall decision is not associated with the true event type. | Chi-square test
H₀3 | The mean SUS score of KALINGA is not significantly different from 68. | One-sample t-test
Reject H₀ when p < 0.05.
Notes: [Researcher 3] | 2:35 to 2:50
H₀1: no difference between KALINGA and the reference; paired t-test. H₀2: the fall decision is not associated with the true event; chi-square. H₀3: the SUS score is not different from 68; one-sample t-test.

## Slide 10
SCOPE AND DELIMITATION
10
Scope
• One 3D-printed pendant and the KALINGA Monitor app
• Temperature, pulse rate, SpO₂ spot checks
• Fall, SOS, no-movement, missed-dose alerts
• Four-slot carousel; placebo tablets in tests
• Volunteers and older adults after IRB approval
Delimitation
• Not a medical device; no diagnosis
• Spot checks only, not continuous ECG or BP
• Needs Wi-Fi to send alerts
• Falls simulated only by researchers on mats
• One day of doses per filling
Table 4. Trials Conducted in the Study
Trial set | Measures | Number | Reference
Health checks | Temperature, PR, SpO₂ | 10 + ____ volunteers | Clinic instruments
Falls | Decision, alert time | ____ falls; ____ activities | Video
Doses | Slot, timing, record | 5 days × 4 slots | Observer log
Delivery, battery, SUS | Messages, hours, SUS | 5 days; ____ users | App log; survey
Notes: [Researcher 3] | 2:50 to 3:05
We cover one pendant and the app, spot checks, four kinds of alerts, and a four-slot carousel. It is not a medical device, and only we simulate falls.

## Slide 11
METHODOLOGY
Research Design
11
This is a quantitative developmental study with method-comparison and usability tests. The pendant is built and calibrated first, then tested against reference instruments and in daily use.
1. Build and calibrate
2. Program logic
3. Run the trials
4. Analyze results
Table 5. Variables of the study
Independent (varied) | Dependent (measured) | Controlled (held fixed)
Method (KALINGA; reference) Event (fall; activity) Dose slot and time | Bias, LoA, Arms Sensitivity, specificity, alert time Correct slot, on time; SUS | Room 26–30 °C; 5-min rest Forehead 3–5 cm; same finger Same Wi-Fi; placebo tablets
Notes: [Researcher 3] | 3:05 to 3:22
This is a quantitative developmental study. We build and calibrate, then test against reference instruments and in five days of daily use.

## Slide 12
METHODOLOGY
Materials and Device
12
Device specification
• Shell: Baymax-inspired, 3D-printed white PLA, about 60 × 83 × 39 mm, breakaway cord.
• Checks: MLX90614 forehead temperature; MAX30102 finger pad for pulse rate and SpO₂.
• Safety: MPU6050 fall and no-movement sensing; SOS button; 10-s cancel window.
• Medicine: 4-slot carousel turned by an SG90 servo; vibration, beeps, OLED message.
• ESP32-C3 with Wi-Fi to Firebase; KALINGA Monitor app; 1000 mAh USB-C battery.
Table 6. Materials and Components of KALINGA
Material / Component | Quantity | Purpose
ESP32-C3 SuperMini | 1 | Controller, Wi-Fi
MLX90614 infrared thermometer | 1 | Forehead temperature
MAX30102 pulse oximeter module | 1 | Pulse rate and SpO₂
MPU6050 accelerometer | 1 | Falls and inactivity
0.91-in OLED display | 1 | Time, doses, results
SG90 micro servo and printed carousel | 1 | Pill dispensing
Vibration motor and buzzer | 1 each | Reminders and alarms
LiPo 1000 mAh, TP4056 USB-C charger | 1 each | Power
White PLA, black paint | 60 g | 3D-printed shell
Reference thermometer and pulse oximeter | 1 each | Reference readings (borrowed)
20-cm gym mats, knee and elbow pads | 2; 1 set | Simulated falls
Notes: [Researcher 3] | 3:22 to 3:40
KALINGA is a 3D-printed, Baymax-inspired pendant with a forehead thermometer, a finger pad, a fall sensor, an SOS button, and a four-slot pill carousel. Parts are estimated at about 2,150 pesos.

## Slide 13
METHODOLOGY
Data Gathering Procedure
13
Step 1
Secure IRB approval, consent
Step 2
Design the shell in CAD
Step 3
3D-print, sand, paint
Step 4
Wire and test electronics
Step 5
Program; set up app
Step 6
Calibrate sensors, fall rule
Step 7
Vital-sign agreement trials
Step 8
Simulated falls, activities
Step 9
Dose, battery, usability tests
Step 10
Bland–Altman, t-tests, χ²
Figure 3. Data gathering procedure
Notes: [Researcher 1] | 3:40 to 4:00
Ten steps: IRB approval; CAD design; 3D printing; wiring; programming and app setup; calibration; agreement trials; simulated falls; dose, battery, and usability tests; and analysis.

## Slide 14
METHODOLOGY
Treatment of Data
14
Formulas
d = KALINGA − reference
difference per paired reading
Bias = Σd ÷ n; LoA = Bias ± 1.96 SD
mean bias, 95% limits of agreement
Arms = √(Σd² ÷ n)
SpO₂ root-mean-square error (≤ 4%)
Sens. = TP ÷ (TP + FN); Spec. = TN ÷ (TN + FP)
falls detected; activities ignored
t_alert = t_app − t_event
alert time (s)
Statistical treatment
SOP | Question | Tool
S1 | Agreement with reference | Bias, LoA, Arms; paired t-test
S2 | Falls and alerts | Sensitivity, specificity; χ²
S3–S4 | Doses; delivery; battery | Percentages; mean ± SD
S5 | Usability | SUS; one-sample t-test vs 68
S6 | Cost | Comparison with Table 1
α = 0.05; software: Microsoft Excel and JASP.
Notes: [Researcher 1] | 4:00 to 4:20
Accuracy is judged by bias, limits of agreement, and Arms; falls by sensitivity and specificity; alerts by time to the app. We use Excel and JASP at alpha 0.05.

## Slide 15
METHODOLOGY
Ethical Considerations and Safety
15
Ethics.
IRB approval first; consent and assent from every participant (ISEF Form 4).
Non-invasive checks only; placebo candy tablets; no real medicine given.
Falls simulated only by researchers on 20-cm mats, with pads and supervision.
Private Firebase account; coded ids; data deleted after the fair (RA 10173).
Risk and safety.
LiPo battery: protected USB-C charger; no puncture or heat.
3D printer and soldering: adult supervision and ventilation.
Neck cord: breakaway clasp to prevent choking or strangling.
Shared sensors: wiped with 70% alcohol between participants.
No data were fabricated; all sources cited in APA 7th edition.
Notes: [Researcher 1] | 4:20 to 4:35
The IRB approves our plan first, and everyone gives consent. We use placebo tablets, simulate falls only on mats, and protect the data under the Data Privacy Act.

## Slide 16
QUESTIONS
& ANSWERS
Notes: [Researcher 3] | 4:35 to 4:42
That ends our presentation. We are now ready for your questions. Thank you.

## Slide 17
THANK YOU
Notes: Leave this slide up after the closing line of slide 17.